Oscilloscopes
A Basic Introduction
Casper Amateur Radio Club – W7VNJ
https://casperarc.net
Tate Belden – KA7O
ka7o@arrl.net
19Feb, 2025 - 1900hrs

---
I’m a ‘tronics hobbyist
Tate, KA7O; an AE, VE – fixer of things
Once a CET at a Commercial Shop
I am NOT an Oscilloscope expert!
© 2025. This work is openly licensed via a Creative Commons BY-NC-ND license. For information, see: https://creativecommons.org/licenses/by-nc-sa/4.0/

---
Copy of this presentation & extras
Find all this, the slide deck and the handout sheet on the club’s website https://casperarc.net/library

---
Online Library Link
On the main page of the web site, find the navigation block on the left, click the link for ‘Online Library’

---
The Online Library
That will take you to this, the Online Library – find the Presentations folder

---
The Presentations Folder

---
Disclaimer and warming
This is all your own risk! If you’re at all unsure, get things clarified before you probe and risk yourself or property
I assume and accept NO Responsibility or Liability for your actions or efforts
It’s all up to you
Seriously, you do this stuff wrong enough and you can not only burn things down, but get dead - fast

---
OK, enough’s enough
Now that’s all outta the way - Let’s get started!

---
The word, “Oscilloscope”
Oscillate (rhythm/swing) and Scope (to see)
See the rhythm, see the swing
O’scope or even just Scope

---
What is an O’scope?
Tool for visualizing signals, waveforms and patterns

---
So, what is it?
A fancy Voltmeter with movies
A Voltmeter takes a snapshot – a single point in time
O’scope shows voltage over time
Allows seeing modulation, frequency and more
If you’re not familiar with using a voltmeter, an O’scope will likely not help you, much

---
 Main Components & Controls

---
Main Components and Controls
1 – Display Screen
This is where you’ll see your signal – may have it’s own ‘comfort’ controls – focus, persistence, brightness, position, etc
It’s an X Y plot, a graph
The horizontal axis X, is time, and the vertical axis Y, is voltage
Two major types of display:
LCD or PC screen on Digital Storage Scopes
CRT on Analog Scopes

---
Other devices with an XY display
VNA - Vector Network Analyzer
SA - Spectrum Analyzer
Signal Analyzer
Vectorscope
(NTSC folks know this one)
Not a comprehensive, nor complete list

---
Side Trip - explanation
1a – DSO and Analog
A digital storage oscilloscope (DSO) samples signals and constructs a display using digital processing, allowing for permanent signal storage and extensive waveform analysis
Analog Scopes trace signals directly on a screen, providing a continuous visual representation of the waveform

---
Main Components and Controls
1b – Scope Types
LCDs or PC displays imply presence of computing power
This offers a LOT of extra tools an analog scope can’t offer, or if they do, they’re expensive
DSOs are VERY good at measuring a LOT of things – enough so, it can become confusing or distracting

---
Main Components and Controls
2 – Input Channels
Most oscilloscopes have two or more input channels, allowing you to compare two (or more) signals at once
Typically BNC connectors on the front
I’ll be focusing on a single channel to keep things simple

---
Main Components and Controls
3 - Time Base (Time/Div)
This controls how much time each division on the horizontal axis represents
If you’re looking at a high-frequency signal, you’ll set it to a small time division

---
Main Components and Controls
3a – Bandwidth and depth
Bandwidth refers to the highest frequency that will pass through to the A/D converter in the scope with reasonable accuracy
For sine waves, bandwidth must equal or exceed the fundamental frequency
For non-sinusoidal wave forms, a bandwidth 5 or more times the fundamental is a good starting point

---
Main Components and Controls
3b – depth
Depth refers to the amount of memory a DSO can allocate for data acquisition, also known as record length – often shared across inputs!
This memory depth is interdependent on the bandwidth and sampling rate to capture and store detailed wave forms

---
Are those weeds?
3c – Sample Rate
Is the speed at which the scope's A/D converter takes measurements of the incoming signal, measured in Samples per Second or Hertz
A higher sample rate allows for a more detailed and accurate representation of the waveform

---
A trip into the weeds
3c – The relationship
The relationship between bandwidth and sample rate is not always straightforward
While the Nyquist* theorem suggests that the sample rate should be at least twice the bandwidth to accurately reproduce a signal, modern DSOs often use techniques like equivalent-time sampling to achieve a bandwidth higher than the sample rate
* https://en.wikipedia.org/wiki/Nyquist%E2%80%93Shannon_sampling_theorem

---
A trip through the weeds
3d – Analog superiority?
Analog scopes just show you what they ‘see’, not what they’ve calculated to see
Bandwidth is limited by their input amplifiers, filters and physics of sweeps
Resolution is limited to the phosphor dot size and your eyes
But, you can’t store a sample to review or look at closer, or with something else

---
Lets get back on track
Yada, yada, yada
The depth and breadth of these kinds of discussions can easily exceed the absurd
Know that wider bandwidths, deeper depths and faster sampling rates all cost more – this stuff can get expensive, fast
Know your tool’s limits – rest is for engineers to hash out after hours

---
Main Components and Controls
4 – Vertical Scale
Back to what we were on about….
This controls the vertical scale, or how much voltage each division represents
If you’re looking at a 5V signal, you might set this to 1 volt per division so that the signal fits nicely on the screen

---
Main Components and Controls
3 and 4 a – Auto Set
Many Scopes, and almost all DSOs have an ‘Auto Set’ or ‘Auto Config’ button
When connected to a signal, press this button and give it a few seconds
It’ll automagickly configure Voltage and Time Base for viewing, but may not show all the details you’re looking for
Use your controls to adjust if needed

---
Main Components and Controls
5 - Probes
These are what you use to connect the oscilloscope to your circuit
They typically have a ground clip and a tip to measure the signal
Ground establishes the ‘zero point’ of a signal measurement

---
Main Components and Controls
5a – Probes & Grounds
Understand the difference between chassis and signal ground
Probes should be connected to chassis ground
When connected to chassis ground, High Current Short Circuits to safety ground can result – destroying equipment

---
Danger, Will Robinson!
Probe, signal and chassis ground

---
Danger, Will Robinson
Probe, signal and chassis ground
When measuring circuits connected to mains power, the ground connection of the oscilloscope can cause issues if not properly managed
The oscilloscope's ground is typically connected to the mains earth safety ground through the AC plug, which can lead to problems if the circuit being measured is also connected to earth ground

---
Danger, Will Robinson
Probe, signal and chassis ground
For safety, especially when testing mains-connected circuits, use an isolation transformer to connect to power
if you are measuring a signal that is not referenced to chassis ground, you can use a differential probe ($$)
Note that the oscilloscope's CAT II rating indicates the maximum voltage the oscilloscope can handle
It does not guarantee your safety

---
Danger, Will Robinson
Probe, signal and chassis ground

---
Main Components and Controls
6 - Trigger
The trigger helps stabilize the waveform so it doesn't jump around (opening slide’s image was badly triggered)
It essentially tells the oscilloscope how and when to start drawing the waveform
Auto-trigger (auto-set) can be your friend

---
Probing and Signal Types

---
Probing and Signal Types
Probing a Circuit
Let’s connect the probe to a signal generator output
Attach the ground clip to the generator’s ground and the tip to the output terminal
Be aware of grounding issues when measuring circuits

---
Probing and Signal Types
Signal Types
You’ll often encounter sine-like waves for analog signals or square-ish waves for digital
Here’s what waves look like on the screen

---
Probing and Signal Types
Analog and Digital signals

---
Probing and Signal Types
Analog and Digital signals
So little is absolute in this world
Even “square waves” will have a rise time, a leading edge slope – if you look close enough
Likewise, a fall time, the trailing edge also has a duration
This is one of the advantages of a DSO, you can store the sampled signal and then ‘Zoom In’ to see how fast rise and fall times may be and their shapes
Special for looking at impedance, or capacative effects, this can be quite handy

---
Demo & Details
1x vs 10x Probes
A 1x probe does not attenuate the signal, making it suitable for measuring low-voltage signals under 1-volt, while a 10x probe reduces the signal amplitude by a factor of 10, which is beneficial for high-voltage and high-frequency measurements.
The 10x probe has a higher bandwidth and lower input capacitance, reducing the load on the circuit under test, but it also reduces the dynamic range and sensitivity of the oscilloscope.
In most cases, a 10x probe is preferred unless you need to measure very small signals with high resolution.

---
Demo & Details
10x Trimming
A 10x probe includes a trimming capacitor that helps equalize the signal and compensate for the low-pass filter effect caused by the series resistance and capacitance in the probe
Trimming a 10x probe involves adjusting this capacitor to ensure that the probe accurately represents the signal being measured

---
 A Demonstration – and details

---
Demo & Details
10x Trim Demo
Watch - I’ll do it up here
Connect probes – ensure 10x is selected
Clip to output on O’scope
   (most Scopes have a 1khz signal source for this)
Adjust trim cap to remove leading spike
Looking for a flat top wave form
Test again with 1x and 10x - verify

---
Measuring Frequency and Voltage

---
Frequency and Voltage
Frequency Measurement
We can measure the frequency of the waveform by counting how many cycles we see in one time division
Or we can push the button on a DSO

---
Frequency and Voltage
Voltage Measurement
To measure voltage, look at the height (amplitude) of the wave form
If each division represents 1V, and our peak-to-peak measurement spans 4 divisions, that’s a 4V peak-to-peak signal
Or we can push the button on a DSO

---
Time for Questions
Rest of this is hands on, but if you’ve questions that’d be good for everyone, ask away!

---
Now, let’s play!
If you’ve your own – cool
If not, come up here and let’s try a few out

---
Oscilloscopes
An Introduction
Casper Amateur Radio Club – W7VNJ
https://casperarc.net
Tate Belden – KA7O
ka7o@arrl.net
© 2025. This work is openly licensed via a Creative Commons BY-NC-ND license. For information, see: https://creativecommons.org/licenses/by-nc-sa/4.0/